SoftPro™ Static-Gel Hip / Knee Orthosis (HKO)

TREATMENT RATIONALE:

Facilitate rehabilitation of the hip/knee in treatment of non-fixed adduction of the hip and flexion of the thigh. A static brace that may be used to maintain end range or to provide incrementally restore range of motion through progressive extension stretching with frequent brace modification.

FUNCTIONAL OBJECTIVES:

Maintain or increase hip Range of Motion, allowing for increased abduction of the hip. Increased abduction improves pericare and functional use of the hip which can significantly be improved by increasing hip abduction. Flexion of the thigh can also be improved by using the OCSI HKO by improving the alignment of the thigh and reducing uncontrolled thigh flexion.

ORTHOTIC TREATMENT:

- Use PROM to passively stretch the affected joint capsule(s), connective tissue, tendons, and muscles.

- Slowly and gently stretch the hip and thigh to the point of noticeable resistance only (no discomfort). Hold for a minute to allow the extension release of the affected joint.

- Determine the HKO abduction bar desired setting to provide the desired orthotic therapy. For a static setting to maintain end range, set the hinges at comfortable end range. For a progressive extension orthotic therapy setting, add approximately 10° of abduction to the comfortable end range. When the device is applied, a gentle stretch will be applied to the affected joint(s).

- Open the lateral cuff flaps.

- Holding the hips at the furthest point of comfortable abduction, position the HKO at mid-thigh.

- Secure the thigh cuff flaps for each thigh cuff.

- Pull the abduction bar setting knob up, and slide the abduction bar to provide the desired amount of abduction.

- Palpate the hip adductors to check tension. The adductors may be taut upon initial brace application. Re-palpate after several minutes. There should be a noticeable softening of the tissue after several minutes of wear.

- Check all cuff flaps for pressure. Two fingers should be able to be inserted between the flap and the skin. Loosen flaps if necessary.

- Determine wearing schedule per therapy evaluation and physician’s order.

- Incrementally increase wearing time per patient tolerance and patient care plan up to a maximum of six hours on per shift. A minimum of three to four hours of wear daily after the adaptation period is recommended for best results.

- The OCSI HKO can be used in the recumbent or seated position. Always check the fit when changing the patient’s position significantly to ensure the device is fitting properly.

- Release and check for skin redness or pressure or patient discomfort every two to three hours at a minimum.  Remove the orthotic device immediately if significant redness or pressure is evident. Notify the appropriate staff member(s) immediately and document any significant redness or signs of adverse pressure or shear. Discontinue device use until the skin integrity issues are resolved, and the device is modified or the wearing schedule is altered to eliminate potential skin integrity problems.

- Follow manufacturer’s instructions for care of the orthotic device. Always inspect the device between applications to ensure the soft goods are properly in place, the device settings have not been altered, and the device has not been soiled or would not provide any other risk to the patient prior to application.

- Check device settings for continued application of the desired amount of abduction at least once a month. Re-adjust the abduction bar setting to maintain 5° to 10( of extension beyond the point of resistance to stretch as needed for progressive extension orthotic therapy.

AIR STARTER KIT (Optional Addition)

TREATMENT RATIONALE:

For severe adduction of the hip (≤ 15°) an Air Starter Kit may be necessary to initiate hip abduction therapy. The minimum setting on the HKO Abduction Bar is 15° of abduction. For severe adduction, the Air Starter Kit replaces the abduction bar of the HKO to allow gentle air bladder Low Load Prolonged Stretch to incrementally abduct the hip to approximately 20° of abduction, at which time the HKO abduction bar should replace the air bladder.

INSTRUCTIONS:

- Remove the abduction bar and rigid thigh cuffs of the HKO by removing the HKO cover.

- Replace the abduction bar and cuffs with the air bladder and cuffs, positioning the air bladder so that the fill nozzle is positioned towards the patient’s feet.

- Apply the HKO as described above.

- Using the clear end of the inflation/deflation bulb, pump air into the air bladder to slowly abduct the hip.

- Evaluate the fit as described above.

- It may be necessary to use the black end of the inflation/deflation bulb to release air from the air bladder prior to removing the device.

- As instructed above, always check for redness, and report any skin integrity issues to the appropriate staff member(s). Discontinue device use until the skin integrity issues are resolved, and the device is modified or the wearing schedule is altered to eliminate potential skin integrity problems.

- Replace the air bladder with the abduction bar after 20° of abduction has been achieved using the air bladder starter kit.

LAUNDRY INSTRUCTIONS:

- Always remove soft cover from frame before washing.

- Close all hook and loop attachments on soft cover and place in enclosed laundry bag.

- Hand or machine wash, gentle cycle with mild detergent. DO NOT USE COMMERCIAL WASHERS OR HOT WATER.

- No bleach or fabric softener.

- Air dry.

WARNING: The product requires a physician’s order and must be properly applied by a trained professional.  The product is designed for single patient use only in order to avoid cross contamination.  Any substitution or removal of the product’s parts voids the manufacturer’s warranty.  OCSI/NeuroFlex, Inc. will assume no liability if the above instructions are not followed.

OCSI/NeuroFlex, Inc.

FL PHONE: (800) 375-0207  ●  FL FAX: (727) 525-1424  ●  CA PHONE: (800) 652-1136  ●  CA FAX: (888) 873-7853